Suspended Solids (Sludge Concentration) Sensor Operation Manual
[image: image41.jpg]



Table of Contents

1Chapter 1 Specifications


2Chapter 2 Product Presentation


22.1 Product Information


32.2 Safety Information


4Chapter 3 Installation


43.1 Installation of Sensors


43.1.1 Quick Dismantling pool side fixed installation


53.1.2 Classic pool side fixed installation


63.1.3 Railing fixed installation


73.2 Connection of Sensor


7Chapter 4 Interface and Operation


74.1 User Interface


74.2 Parameter Setting


8Chapter 5 Calibration of Sensor


85.1Factor calibration


105.2Four-point calibration


175.3Curve calibration


21Chapter 6 Communication Protocol


27Chapter 7 Maintenance


277.1 Sensor Cleaning


287.2 Inspection on the Damage of Sensor


287.3 Sensor Blade Replacement


28Chapter 8 Special Description





Chapter 1 Specifications

	Specifications
	Details

	Measurement Range
	0.01-20000 mg/L、0.01-50000 mg/L 、0.01-120000 mg/L

	Measurement Accuracy
	Less than ± 5% of the measured value (depending on sludge homogeneity)

	Repeatability
	±2%

	Resolution Ratio
	0.01~1 mg/L，It depends on different ranges

	Pressure Range
	≤0.4Mpa

	Flow velocity
	≤2.5m/s、8.2ft/s

	Material
	Main Body: SUS316L/Titanium Alloy /PVC
Upper and Lower Cover：POM；Cable: PUR

	Power Supply
	12VDC

	Communication Protocol
	MODBUS  RS485

	Operating Temperature
	0~45℃（not freeze）

	Weight
	1.65 KG

	Level of Protection
	IP68/NEMA6P

	Cable Length
	Standard 10-Meter Cable, Max Length: 100 Meters


Table 1 Technical Specifications of Suspended Solids (Sludge Concentration) Sensor
Note: The specifications of the product are subject to change without prior notice.

Chapter 2 Product Presentation
2.1 Product Information
Suspended solids (Sludge Concentration) sensor is based on the method of combination of infrared absorption scattered light. Infrared light which the photosource sends out scattered by suspended particles in the sample. And finally converted to electrical signals by photoelectric detector and getting the suspended solids concentration of the sample after treatment by analog and digital signals.

The product is widely used in sewage plant, water plant, water station, surface water, farming, industry and other fields. The size of the suspended solids (Sludge Concentration) sensor is shown in Figure 1.

SUS316L/Titanium Alloy External Dimension：

No automatic cleaning device:
[image: image2.png]254
193.5

\\\\\\\\\\\\\\\\\\\\1





With automatic cleaning device:
[image: image3.png]



PVC External Dimension：

No automatic cleaning device:
[image: image4.png]250.5

R |/ E—

AAERAREIESSERRANIN

oY




With automatic cleaning device:
[image: image5.png]ARSI

oY




Figure 1 Dimension Diagram of suspended solids （Sludge Concentration） Sensor
2.2 Safety Information
Please read this manual completely before opening the package, installing or using. Otherwise it may cause personal injury to the operator, or cause damage to equipment.

Warning labels

Please read all labels and signs on the instrument, and comply with the security label instructions, otherwise it may cause personal injury or equipment damage.

[image: image1.png]


[image: image40.jpg]


When this symbol appears in the instrument, please refer to the operation or safety information in the reference manual.

While this symbol indicates an electric shock or risk of death from electric shock.

Please read this manual completely. Pay particular attention to some notes or warnings, etc. To ensure that the protective measures provided by the equipment are not destroyed.
Chapter 3 Installation
3.1 Installation of Sensors

3.1.1 Quick Dismantling pool side fixed installation

[image: image6.png](BZ]-03

14 bolt HE*12

11 |rotary union(M6 to B )
10 H8 hex nut

9 W8 flat gasket

8 I8 spring washer
7 W8 flat gasket

] pipe c¢lamp

3 L type pipe hoop
4 sus304 pipes (638mm)
3 end cap

9 Vounting plate

1 W8 expansion bolt

num.

Part

11

&

e

12





Figure 2 Quick Dismantling pool side installation sketch map
3.1.2 Classic pool side fixed installation

[image: image7.png](BJD-03

bolt  H6*12

rotary union{(M6 to R} )

o | —a| oo

H8 hex nut

bolt W8%16

sus301 pipes (#38mm)

end cap

pipe hoop (#10)

[ O [RFE IS

W4 expansion bolt

num.

Part

-

«





Figure 3 Classic pool side fixed installation sketch map
3.1.3 Railing fixed installation

[image: image8.png]L67]-03

14 boll H6x12

13 8 spring wesher

12 W8 flat gasket

11 Younting plate

10 bolt 16#120

9 [ type pipe hoop

8 | rotery wnion(¥6 o B3 )

7 L type pipe hoop

[} pipe clamp

3 sus30 pipes (938mm)

1 ond cap

3 W8 kex nul

2 W8 Ial gasket

1 W8 hex nut
num, Part





Figure 4 Railing fixed installation  sketch map

3.2 Connection of Sensor
The sensor should be correctly connected by the following definition of wire core:

	Serial No.
	1
	2
	3
	4
	5

	Sensor Cable
	Brown
	Black
	Blue
	White
	Yellow + Green

	Signal
	+12VDC
	AGND
	RS485 A
	RS485 B
	Ground lead


Chapter 4 Interface and Operation

4.1 User Interface
The sensor is connected to the computer using RS485 to USB, and then use Modbus Poll to connect

Note: Modbus Poll software is a general software that can be downloaded online.
4.2 Parameter Setting

1、Click “Setup” on the menu bar, select “Read / Write Definition”, and then set the parameters（The slave address for the first time is the slave label）, then enter “30” for Quantity in the dialog box , click “OK” 
[image: image9.png]Read/Write Definition X
Savein: |1 oK
Function: |03 Read Holding Registers (4x) v Cancel
Address: |0 Protocol address. E.g. 40011 - 10
Quantty: |30
Scan Rate: |1000 [ms] Apply
Disable
[[]Readwrite Disabled
[[] Disable on emor Readjwiite Once
View

Rows
@10 O20 Oso O100 O FittoQuantity

[[]Hide Alias Columns [C]PLC Addresses (Base 1)
[JAddress in Cell [] Enron/Daniel Mode




Note: After the slave address is changed, the new address will be used for communication and the slave address for the next time connection is also the most recently changed address.
2、Click “Connection” on the menu bar, select the first line in the drop-down menu “Connection setup”（The baud rate for the first time is the slave label）and click “OK”.

[image: image10.png]Connection Setup

Connection
Serial Port

Serial Settings
Prolific USB-to-Serial Comm Port (COM4)

9600 Baud

8 Data bits

None Parity v

1

Stop Bit

Remote Modbus Server

or Node

Name

Advanced

OK

Cancel

Mode

@RTU  OAsCI

Response Timeout
[1000 Jims]

Delay Between Polls
[1000 Jims]




Note: Port is set according to the Port number of the connection

Note: If the sensor has been connected as described, and “Timeout Error” appears on the software “Display status”, it means that the connection is failed; remove and replace the USB port or check the USB to RS485 converter, repeat the above procedure until the sensor connection is successful.
Chapter 5 Calibration of Sensor

There are 2 kinds of calibration methods for the sensor, but only one type is valid. Which calibration mode is used depends on the actual needs of the site. For the selection method, see the communication protocol.
The specific calibration can be carried out according to the following steps.
Note: Curve calibration is used as the base calibration and coexists with the other 2 calibration mode.
5.1Factor calibration
Calibration of suspended solids (sludge concentration) requires the use of suspended solids standards solution. During the process of calibration, make sure the lens of probe is 15cm far from the bottom of calibration cup and no bubbles in front of the lens, meanwhile, it is recommended to keep the correction cup away from the light. Specific steps are as follows
If there is a large deviation between the measured value and the standard value, the slope of calibration curve needs factor correction.

1、Connect the sensor to the Modbus software;
2、Set the relevant parameters and wipe the sensor;
3、Select “06” in the menu bar to enter “27” for Address in the dialog box, “1” for Value, then click “Send”, as shown below;
[image: image11.png]Write Single Register

Cancel

Slave D,

Address:

Value:

Result
Response ok

[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




4、Select “16” in the menu bar to enter “06” for Address in the dialog box, “2” for Quantity and change type into “Float CD AB”. Double-click the up on the right to enter “1” for the value. Click “OK”, then click “Send”, as shown below;
[image: image12.png]16: Write Multiple Registers

Saein |l | [ Send
Address ‘06 ‘ Cancel
Quantty:  [2 | Edit
Type [Float cD AB v 5
Save





5、Slowly immerse the sensor into the suspended solids (sludge concentration) standard solution;
6、Wait for the value to be stable and record the stable value;
7、Calculate the correction factor. The correction factor equals to the standard solution value divided by the value measured in the 6 step. (Factor = standard solution value / stable value);

8、Select “16” in the menu bar to enter “06” for Address in the dialog box, “2” for Quantity and change type into "Float CD AB". Double-click the up on the right to enter “Factor Value” for the value(The factor value is the calculated value in step 7). Click “OK”, then click “Send”.

Assume that the factor value calculated in step 7 is 0.93, as shown in the figure below:

[image: image13.png]16: Write Multiple Registers

Save ID: |1 oos-093 ] Send

Address: |06 | cancel |
Quantity: |2 [ Eat
Type: |FoatcpaB v o




[image: image14.png]Tx =6469: Err =2204: D = 1: F = 03: SR = 1000ms

BRI CNMNEE)

ias| 00000
0

Suspended solids value 5667.26

Suspended solids factor _

0

Alias 00010 Alias 00020

0 Manual Brushing Order 0|

Brushing Time 1 Automatic Brushing Order 0
Response Time 1 0|
Suspended solids 2 0|
Probe Humidity | 0 0|
1 0

Probe Baud Rate 9600 0
Probe Slave Address 1 0
Serial No. 1 221 0|
Serial No. 2 8329 0





5.2Four-point calibration

Four-point calibration of suspended solids (sludge concentration) requires the use of suspended solids standards solution. The specific steps are as follows:

1、Connect the sensor to the Modbus software;

2、Prepare the four suspension standards required for four-point calibration. Generally, they are 0 point, 0.25 times or 0.5 times or full scales point of the range, wipe the sensor clean;
3、Refer to steps 3-4 of 5.1 section, set the calibration mode as factor, the factor has a value of 1, in this way, the measured values of each standard solution are measurement value.

4、Slowly immerse the sensor into the first suspension standard solution, record the first standard solution value (standard solution value is target value) and standard solution measurement value (standard solution measurement value is actual value), clean and wipe clean. Slowly immerse the sensor into the second suspension standard solution. record the second standard solution value and standard solution measurement value, clean and wipe clean. Slowly immerse the sensor into the third suspension standard solution, record the third standard solution value and standard solution measurement value, clean and wipe clean. Slowly immerse the sensor into the fourth suspension standard solution, record the fourth standard value and standard solution measurement value, clean and wipe clean; (Note: standard value divided by measurement value < 2)

5、Select “06” in the menu bar to enter “27” for Address and “3” for Value, and then click “Send”, as shown below;

[image: image15.png]Write Single Register

Slave D,
Address:

Value:

Result
Response ok

[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




6、Select “06” in the menu bar to enter “28” for Address and “1” for Value in the dialog, and then click “Send”, as shown below;

[image: image16.png]Write Single Register

Slave D,

Address:

Value:

Result
Response ok

[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




7、Select “16” in the menu bar to enter “20” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter Value as “first standard solution value”, click “OK”, then click “Send”, as shown below;
[image: image17.png]16: Write Multiple Registers

Save ID: |1 o2o-217 ] Send
Address: |20 | cancel |
Quantity: |2 [ Eat
Type: Float CD AB P




8、Select “16” in the menu bar to enter “22” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter Value as “first standard solution measurement value”, click “OK”, then click “Send”, as shown below;
[image: image18.png]16: Write Multiple Registers

Slave ID: |1 Jo2-267 | Send
Address: |22 Cancel
Quantity: |2 Edit
Type: Float CD AB v =
Save





9、Select “06” in the menu bar to enter “28” for Address and “2” for Value, and then click “Send”, as shown below;

[image: image19.png]Write Single Register

Slave D,

Address:

Value:

Result
Response ok

[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




10、Select “16” in the menu bar to enter “20” for Address in the dialog box, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “second standard solution value”, click “OK”, then click “Send”, as shown below;

[image: image20.png]16: Write Multiple Registers

Saein |l |
Address: |20 |
Quantty.  [2 |
Type [Float cD AB v

Send

Cancel

Edit

Qpen

Save





11、Select “16” in the menu bar to enter “22” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “second standard solution measurement value”, click “OK”, then click “Send”, as shown below;

[image: image21.png]16: Write Multiple Registers

Slave ID: |1

Address: |22

Quantity: |2

Type: Float CD AB





12、Select “06” in the menu bar to enter “28” for Address, “3” for Value, and then click “Send”, as shown below.

[image: image22.png]Write Single Register

Slave D, end
Address: Cancel
Value:

Result

Response ok
[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




13、Select “16” in the menu bar to enter “20” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “third standard solution value”, click “OK”, then click “Send”, as shown below;

[image: image23.png]16: Write Multiple Registers

Slave ID: |1 o0-2100 | Send
Address: |20 Cancel
Quantity: |2 Edit
Type: Float CD AB v =
Save





14、Select “16” in the menu bar to enter “22” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “third standard solution measurement value”, click “OK”, then click “Send”, as shown below;

[image: image24.png]16: Write Multiple Registers

Slave ID:
Address:
Quantity

Type:

22

Float CD AB





15、Select “06” in the menu bar to enter “28” for Address, “4” for Value, and then click “Send”, as shown below.

[image: image25.png]Write Single Register

Slave D,
Address:

Value:

Result
Response ok

[ Close dislog on "Response ok

Use Function
® 06: Wite single register
(O 16: Wi muliple registers




16、Select “16” in the menu bar to enter “20” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “fourth standard solution value”, click “OK”, then click “Send”, as shown below;

[image: image26.png]16: Write Multiple Registers

Slave ID: | Send
Address: |20 Cancel
Quantity: |2 Edit
Type Float CD AB v e
Save





17、Select “16” in the menu bar to enter “22” for Address, “2” for Quantity, and “Float CD AB” for Type. Double-click the value that pops up on the right to enter “fourth standard solution measurement value”, click “OK”, then click “Send”, as shown below;

[image: image27.png]16: Write Multiple Registers

Slave ID: | Send
Address: |22 Cancel
Quantity: |2 Edit
Type Float CD AB v e
Save





18、After the calibration is complete, remove the sensor and clean it.

5.3Curve calibration
Suspension (sludge concentration) calibration requires the use of turbidity standards solution. The specific steps are as follows:
(If you enter the curve calibration, you must perform the calibration action until step 11, otherwise the sensor will be in the calibration mode all the time. If the normal measurement cannot be performed, you can choose to restart after power failure or perform step 12).
1、Connect the sensor to the Modbus software;
2、After setting the parameters according to Section 4.2, select the second column and right click and select “Format”, click “Float CD AB” and wipe the sensor;

3、Select “16” in the menu bar to enter “06” for Address in the dialog box, “2” for Quantity, and change type into “Float CD AB”. Double-click the up on the right to enter “1” for the value. Click "OK", then “Send” as shown below;

[image: image28.png]16: Write Multiple Registers

Saein |l |
Address: |06 |
Quantty.  [2 |
Type [Float cD AB v

Send

Cancel
Edit
Open

Save





4、Select “06” in the menu bar to enter “13” for Address in the dialog box, “1” for Value, click “Send”. The probe type becomes turbidity, as shown below;

[image: image29.png]Write Single Register

Saveid: |1 | [ send
Address \‘3 \ Cancel
Value [ |

Result

Response ok

[]close dialog on "Response ok"

Use Function
(®) 06: Write single register
(O 16: Write multiple registers




[image: image30.png]O |® N v s wnv|=|o

Alias

Turbidity Value

Turbidity Factor

Turbidity Deviation Value

Alias 00010 Alias

0 Manual Brushing Order
Brushing Time 1
Response Time 1
Probe }i\.m.d.ty 0
1
Probe Baud Rate 9600
Probe Slave Address 1
Serial No. 1 221
Serial No. 2 8329




5、Select “16” in the menu bar to enter “04” for Address in the dialog box, “2” for Quantity, and change type into “Float CD AB”. Double-click the up on the right to enter “1” for the value. Click “OK”, then “Send” to start calibration.

[image: image31.png]16: Write Multiple Registers

Saein |l |
Address: |04 |
Quantty.  [2 |
Type [Float cD AB v

Send

Cancel

Edit
Open

Save





6、Enter calibration mode, select “06” in the menu bar to enter “59” for Address in the dialog box, enter “66” for Value, and click “Send”;

[image: image32.png]Write Single Register

Saveid: |1 | [ send
Address \59 \ Cancel
Value \66 \

Result

Response ok

[]close dialog on "Response ok"

Use Function
(®) 06: Write single register
(O 16: Write multiple registers




7、Place the sensor in distilled water. After a while, select “16” in the menu bar to enter “59” for Address in the dialog box, enter “1” for Value;
8、Wait for the value in the box below the figure to be less than 17 and stabilize for a period of time in address “22”, select “06” in the menu bar to enter “59” for Address in the dialog box, enter “2” for Value click “Send”, close the dialog box;

[image: image33.png]Write Single Register

Slave ID: ! Send
Address: 59 Cancel
Value: &l

Result

NA

[[Jclose dialog on "Response ok”

Use Function
(®) 06: Write single register
(O 16: Write multiple registers




[image: image34.png]olficol| N on [l EIICIIIENIOI)

Alias 00000 Alias| 00010| Alias 00020

Turbidity Value 0 0 Manual Brushing Order 343.043

- Brushing Time 1 -

0 ReponseTime 1

- Turbidity 1 --

Turbidity Factor 1 Probe Humidity 0 0

- 1 -

0 Probe Baud Rate 9600 0

-~ Probe Slave Address 1 -

Turbidity Deviation Value 0 Serial No. 1 221 0
Serial No. 2 8329





9、Select “16” in the menu bar to enter “30”for Address , and “2” for Quantity. Change the Type to “Float CD AB” and double-click the value popped on the right to enter the value into the “known standard value (500-1000NTU)”, click “OK”, then click “Send”;
10、Put the sensor into the standard solution. After waiting for 10s, select “06” in the menu bar to enter “59” and “3”for value click “Send”. After a period of time, wait until the value and input in the box below the value is close and stable, click “Send”. Select “06” in the menu bar to enter “59” then change the Value to “4”, click “Send”, the calibration is completed;

[image: image35.png]Write Single Register

Slave ID: ! Send
Address: 59 Cancel
Value: 4

Result

NA

[Jciose dialog on "Response ok"

Use Function
(®) 06: Write single register
(O 16: Write muttiple registers




[image: image36.png]oo N|o|un|sr|lwv|= of

Alias 00000

Turbidity Value
o

Turbidity Factor -1-
o

Turbidity Deviation Value -(;

Alias| 00010 Alias|

0 Manual Brushing Order
Brushing Time 1
Response Time 1
Turbidity 1
Probe Humidity 0
1
Probe Baud Rate 9600
Probe Slave Address 1
Serial No. 1 221

Serial No. 2 8329




11、After the calibration is completed, select “06” in the menu bar to enter the “13” and “2” for value, and click “Send”. The probe type becomes suspended solids (sludge concentration), as shown below;

[image: image37.png]Write Single Register

Slave D! ! Send
Address 18 Cancel
Value: 2

Result

Response ok

[[]close dialog on "Response ok"

Use Function
(®) 06: White single register
(O 16: White multiple registers




12、If the customer wants to quit the calibration during the calibration or for other reasons, select “06” in the menu bar to enter “59” and Value to enter “33”.

Note: If the curve calibration is inaccurate, it is recommended to use factor calibration to calibrate again.

Chapter 6 Communication Protocol
The sensor is equipped with MODBUS RS485 communication function, please refer to this manual section 3.2 to check the communication wiring. The specific MODBUS RTU table is shown in the following table.

	MODBUS-RTU

	Baud Rate
	4800/9600/19200/38400

	Data Bits
	8 bit

	Parity Check
	no

	Stop Bit
	1bit


	Register Name
	Address Location
	Read/Write
	Data Type
	Register Number
	Descriptions

	Suspended solids /sludge concentration Value
	2
	OR
	Float
	2
	0-Range

	Suspended solids/sludge concentration Factor
	6
	RW
	Float
	2
	0.1-10

	Brushing Time
	11
	OR
	Int
	1
	15/30/60/240/720/1440/4320/ 10080min

	Response Time
	12
	RW
	Int
	1
	1-60s

	Suspended solids/sludge concentration
	13
	RW
	Int
	2
	It should be 2, otherwise, it should be changed to 2.

	Probe Humidity
	14
	OR
	Int
	1
	It is recommended to be less than 10（Greater than 10 indicates that the sensor may have been flooded）

	Probe Baud Rate
	16
	RW
	Int
	1
	W       R

0 stands for 4800      4800

1 stands for 9600      9600

2 stands for 19200     19200

3 stands for 38400     38400

	Probe Slave Address
	17
	RW
	Int
	1
	1-254

	Serial No. 1
	18
	OR
	Int
	1
	First 4 numbers of Serial No.

	Serial No. 2
	19
	OR
	Int
	1
	Last 4 numbers of Serial No.

	Manual Brushing Order
	20
	W
	Int
	1
	Send 66

	Automatic Brushing Order
	21
	W
	Int
	1
	Time between sending intervals: 15/30/60/240/720/1440/4320/ 10080min

	Calibration Mode

	Factor Calibration （use of suspended solids standards solution）

	Step 1
	27
	W
	Int
	1
	Send 1（1 stands for the sensor enable factor correction mode）

	Four-point Calibration（use of suspended solids standards solution）

	Step 1
	27
	W
	Int
	1
	Send 3(3 stands for the sensor enable four-point calibration mode)

	the First Point Calibration

	Step 1
	28
	W
	Int
	1
	Send 1(1 stands for the First Point)

	Step 2 Set Target Value
	20
	W
	Float
	2
	Send Target Value

	Step 3 Set Actual Value
	22
	W
	Float
	2
	Send Actual Value

	the Second Point Calibration

	Step 1
	28
	W
	Int
	1
	Send 2(2 stands for the Second Point)

	Step 2 Set Target Value
	20
	W
	Float
	2
	Send Target Value

	Step 3 Set Actual Value
	22
	W
	Float
	2
	Send Actual Value

	the Third Point Calibration

	Step 1
	28
	W
	Int
	1
	Send 3(3 stands for the Third Point)

	Step 2 Set Target Value
	20
	W
	Float
	2
	Send Target Value

	Step 3 Set Actual Value
	22
	W
	Float
	2
	Send Actual Value

	the Fourth Point Calibration

	Step 1
	28
	W
	Int
	1
	Send 4(4 stands for the Fourth Point)

	Step 2 Set Target Value
	20
	W
	Float
	2
	Send Target Value

	Step 3 Set Actual Value
	22
	W
	Float
	2
	Send Actual Value

	Curve Calibration （use of turbidity standards solution）

	Step 1
	13
	W
	Int
	1
	Send 1 (1 means switch to the turbidity sensor status, ready for calibration)

	Step 2
	59
	W
	Int
	1
	Send 66 (66 means enter calibration mode)

	Step 3
	59
	W
	Int
	1
	Send 1 (1 means ready to calibrate 0 points)

	Step 4
	59
	W
	Int
	1
	Send 2 (2 means calibration 0 point)

	Step 5
	30
	W
	Float
	2
	Write the 2nd point standard value (500-1000NTU)

	Step 6
	59
	W
	Int
	1
	Send 3 (3 means ready to calibrate point 2)

	Step 7
	59
	W
	Int
	1
	Send 4 (4 means calibration point 2)

	Step 8
	13
	W
	Int
	2
	Send 2 (2 means switch to the suspended solids (sludge concentration) sensor status, enter normal measurement)

	In-process cancellation
	59
	W
	Int
	1
	Send 33 （33 to exit calibration mode）


485 analysis：

1、Read the suspended solids (sludge concentration) value
	Register Name
	Address Location
	Read/Write
	Data Type
	Register Number
	Descriptions

	Suspended solids /sludge concentration Value
	2
	OR
	Float
	2
	0-Range


Send the command：01 03 00 02 00 02 65 CB
The equipment return：01 03 04 00 00 40 E0 CA 7B
Send command parsing：

01：device address 01

03：Function code 03 for reading register content
00 02：The starting register address read is 0002

00 02：Read 2 registers
65 CB：CRC16 check code
The device returns the analysis:
01：device address 01

03：Function code 03 for reading register content 

04：The length of the returned data is 4 bytes
00 00 40 E0：The suspended solids (sludge concentration) value read is 7.00（analyze 40 E0 00 00 using IEEE 754）

CA 7B：CRC16 check code
2、Read the brushing time
	Register Name
	Address Location
	Read/Write
	Data Type
	Register Number
	Descriptions

	Brushing Time
	11
	OR
	Int
	1
	15/30/60/240/720/1440/4320/ 10080min


Send the command：01 03 00 0B 00 01 F5 C8

The equipment return：01 03 02 00 0A B8 44

Send command parsing：

01：device address 01

03：Function code 03 for reading register content
00 0B：The starting register address read is 0011

00 01：Read 1 registers
F5 C8：CRC16 check code
The device returns the analysis：

01：device address 01

03：Function code 03 for reading register content
02：The length of the returned data is 2 bytes
00 0A：The brushing time read is 10
B8 44：CRC16 check code
3、Set the manual brush
	Register Name
	Address Location
	Read/Write
	Data Type
	Register Number
	Descriptions

	Manual Brushing Order
	20
	W
	Int
	1
	Send 66


Send the command： 01 06 00 14 00 42 49 FF

The equipment return：01 06 00 14 00 42 49 FF
Send command parsing：

01：device address 01

06：Function code 06 for writing register content
00 14：The register address of write data is 00 20

00 42：Write data content of 66
49 FF：CRC16 check code
The device returns the analysis：

01：device address 01

06：Function code 06 for reading register content
00 14：The register address of the return write data is 20
00 42：Returns modified data content of 66
49 FF：CRC16 check code
4、Set the suspended solids (sludge concentration) factor
	Register Name
	Address Location
	Read/Write
	Data Type
	Register Number
	Descriptions

	Suspended solids/sludge concentration Factor
	6
	RW
	Float
	2
	0.1-10


Send the command：01 10 00 06 00 02 04 00 00 3F 80 63 D5
The equipment return：01 10 00 06 00 02 A1 C9
Send command parsing：

01：device address 01

10：Function code 16 for writing register content
00 06：The starting register address write is 00 06

00 02：Write 2 registers
04：The length data is 4 bytes
00 00 3F 80：The suspended solids (sludge concentration) value write is 1.00（analyze 3F 80 00 00 using IEEE 754）

63 D5：CRC16 check code
The device returns the analysis：

01：device address 01

10：Function code 16 for writing register content
00 06：The starting register address of the return write data is 00 06

00 02：Returns 2 registers
A1 C9：CRC16 check code
Chapter 7 Maintenance

In order to obtain the best measurement results, it is very necessary to maintain the sensor regularly. Maintenance mainly includes cleaning, inspecting damage of the sensor. You can also view the sensor’s status during maintenance and inspection.
7.1 Sensor Cleaning
Both the two lenses on the sensor need to be cleaned and maintained regularly to ensure the accuracy of the measurement on the basis of actual use.  Wash with clean water, then wipe with a cleanser and rag to remove stubborn stains.
7.2 Inspection on the Damage of Sensor

Check the appearance of the sensor to see whether there is damage, if it’s damaged, please contact to the after-sales service center in time for replacement to prevent malfunction of sensor caused by water due to the damage. Note: It is recommended to replace the sealing ring once a year.
7.3 Sensor Blade Replacement

It is recommended that the blade of the sensor should be replaced quarterly with a new rubber one, the specific steps are as follows:
SUS316L/Titanium Alloy：                PVC：
[image: image38.png]Blade



     [image: image39.png]Blade




1. The position of the blade is shown in the figure;
2. Remove the rubber sheet from the scraper;
3. Then apply lubricant to the bracket
4. Load new rubber sheet.
Chapter 8 Special Description

The optional self-cleaning brush is motor-driven. In order to avoid the internal motor gear, please don't move the brush with hand.  The warranty doesn't include the misfunction/damage caused by manually moving of the brush.
Version Number：R1010_200430/R1017_211015/R2010_20220516/R1021_220601









